Standard Operating Procedures

Laboratory Specific

Electrical Safety, Hazardous Voltage, and High Voltage

| Department: | Chemistry & Biochemistry |
| --- | --- |
| Date SOP was written: | 1/15/2013 |
| Date SOP was approved by PI/lab supervisor: | 1/15/2013 |
| Principal Investigator: | Irene Chen |
| Internal Lab Safety Coordinator/Lab Manager: | Irene Chen |
| Lab Phone: | 805-893-8085 |
| Office Phone: | 805-893-8364 |
| Emergency Contact: | Irene Chen, 617-710-8741 (cell) |
|  | (Name and Phone Number) |
| Location(s) covered by this SOP: | Chemistry 1142 |
|  | (Building/Room Number) |

Type of SOP:         X   Process            Hazardous Chemical             Hazardous Class

Purpose

- All electrical work should be done with the system powered off, and then disconnected (unplugged) from the power source.

- Electrical work is defined as any manipulations (electrical or not):

  - inside the casing of electrical equipment.

  - of the inside of a muffle furnace, or any other device with exposed wiring.

  - involving forming or modifications of electrical connections.

Do not modify the inside of a muffle furnace with

your hands or metal tweezers while it is plugged in!

- Soldering is a specific type of electrical work with extra precautions (see #2 and #5).

- Electrical safety should be considered any time when using equipment, whether commercial or not, whether fully insulated or not, whether shielded behind grounded casing or not.

High Voltages (>600V) requires special training, PPE, and precautions.

Physical & Chemical Properties/Definition of Chemical Group

| Health |  | Flammability |  | Reactivity |  |
| --- | --- | --- | --- | --- | --- |

| Chemical Name | CAS | Starting Concentration (percent) | NFPA Ratings (combined solder fumes) |  | Other Hazards Carcinogen, Skin Absorption, etc. |
| --- | --- | --- | --- | --- | --- |
|  |  |  | 3 |  |  |
| Lead | 7440-92-1 | 0-100 |  | Do not breathe |  |
| Tin | 7440-31-5 | 0-100 | 0 | Do not breathe (stannosis) |  |
| Antimony | 7440-36-0 | <2 |  |  |  |
| Copper | 7440-50-8 | <2 | 0 |  |  |
| Rosin | 8050-09-7 | 1-3 |  |  |  |

Potential Hazards/Toxicity

General notes

- Serious shock:

  - More than: 30 V RMS, 42.4 V peak, 50 V DC

  - 10-75 mA

  - More than 10 J

- Lethal shock:

  - More than 375 V

  - More than 75 mA

  - More than 50 J

- > 50 V = “Hazardous Voltage”

  - Special training recommended (including electrical lockout and tagging procedures), but not offered on-site

- >600 V = “High Voltage”

  - Special training required (not offered at UCLA; need to seek out appropriate source)

  - Special PPE required (See PPE section)

  - Check ratings on all test equipment

- AC Risk

  - 60Hz (wall outlet): extremely dangerous; respiratory and cardiac arrest at low current (figure)

  - Higher frequencies are increasingly less dangerous (up to ~50 kHz)

  - Above 100-200 kHz serious burns can occur

  - Above 1MHz, a path to ground and direct contact are not needed for burns to occur

- Avoid being an electrical conduit

  - Be sure to check that the power is off and the equipment is additionally unplugged EVERY time before working on circuitry

  - Discharge capacitors using an inductor before touching or soldering across them

  - Touch circuits with the back of the hand first

  - Wear rubber-bottomed shoes

  - Keep one hand in your pocket when the voltage is on

  - Do not wear metal jewelry or loose metal items

- Keep an open working space of >30 inches from the equipment (>4 ft if there is equipment/grounded surfaces on the other side of the space)

- Work on insulating surfaces

- Do not do electrical work or operate high voltage equipment while tired

- Do not work alone

- If you must probe live wires, use small tips and use electrical tape along all but the last 1/16” of the probes to avoid shorts.  Clip the ground so that you only need to use one hand.

- Use an isolation transformer if there is any chance of contacting line-connected power (AKA the wall outlets)

- Only change fuses when the power is off and the device disconnected.  When changing fuses in the Variacs, make sure to insert the fuse into the loose end first, and then screw the loose end into the Variac.  For other devices, make sure that contact is made with the line side last (and not while in contact with your fingers).

- Always test (with a multimeter) that no accessible parts have been accidently shorted to a live wire before turning equipment back on

  - DC leakage:  Resistance should measure >1M Ω between the line end of the plug and any exposed surface.

- Soldering:

  - Do not breathe in metal fumes, especially if the solder contains lead or tin (use a hood).

  - Always keep the iron in the stand when you put it down.

  - Wash your hands and under fingernails after soldering / handling solder (especially with lead-based solders).

  - Don’t burn yourself.  Do not touch newly soldered joints, or any wires leading away from them.  Hold wires together with a clamp, not your fingers.

  - Discharge all capacitors with a resistor before soldering across them.

Personal Protective Equipment (PPE)

- Eye protection (glasses/goggles)

- Fire-resistant lab coat

- Rubber-soled shoes

- No dangling jewelry/objects ; long hair tied back

- >600V

  - Rubber-soled work boots

  - Rubber, non-conductive gloves

  - Rubber, non-conductive apron

  - Face shield

  - Auxiliary Insulating and shielding mats and guards as appropriate

Engineering Controls

- Elephant Trunk Hoods (for soldering) —Located in 4110

First Aid Procedures

Electrical shock

- If a victim receives a shock that does not cause collapse (respiratory or cardiac arrest), turn off the equipment, and monitor the person for 15 minutes.  Check for and treat any burns.

- If a lab member collapses, the nearest lab member should

  - Check that the collapsed member is no longer connected to the source of electricity.  If so, turn the power off on the system, and unplug the system.  If this cannot be safely done, turn the power off for the lab using the circuit breakers, or use an insulating item (ex: wooden plank) to separate the victim from the source.  If a second lab member is present:  call 911 using a campus phone.

  - When the fallen lab member can be safely approached, check for pulse and breathing, and then call 911 using a campus phone.  Answer their questions.  Do not hang up.

  - If there is no pulse or breathing, qualified individuals should perform CPR or rescue breathing until the person resumes breathing or EMTs arrive.  DO NOT perform CPR on a person with a heart-beat, or rescue breathing on a person who is breathing, even if they are struggling to do so.  Monitor them until help arrives.

- Seek medical help if symptoms persist

Special Handling and Storage Requirements

Soldering

The easiest to use solders are lead based.  Avoid breathing in the soldering fumes.

Spill and Accident Procedure

Electrical fire

- Pull the fire alarm and call 911 from a campus phone in the case of any fire.

- If comfortable, and not in danger, an electrical fire can be extinguished using a class C fire extinguisher .  Make sure to always stand between the fire and an escape route when using a fire extinguisher.

Chemical Spill Dial 911 and call EH&S at 893-3194

Spill – Help contaminated or injured persons.  Evacuate the spill area.  Avoid breathing vapors.  Eliminate sources of ignition if the chemical is flammable.  If possible, confine the spill to a small area using a spill kit or absorbent material. Keep others from entering contaminated area (e.g., use caution tape, barriers, etc.).

Small (<1 L) – If you have training, you may assist in the clean-up effort.  Use appropriate personal protective equipment and clean-up material for chemical spilled.  Double bag spill waste in clear plastic bags, label and take to the next chemical waste pick-up.

Large (>1 L) – Dial 911 (or 310-825-1491 from cell phone) and EH&S at 893-3194 for assistance.

Chemical Spill on Body or Clothes – Remove clothing and rinse body thoroughly in emergency shower for at least 15 minutes.  Seek medical attention. Notify supervisor and EH&S at 893-3194 immediately.

Chemical Splash Into Eyes – Immediately rinse eyeball and inner surface of eyelid with water for 15 minutes by forcibly holding the eye open.  Seek medical attention. Notify supervisor and EH&S at 893-3194 immediately.

# Medical Emergency Dial 9-911

Life Threatening Emergency: Dial 9-911 Contact PI and EH&S as soon as practical. Note: All serious injuries must be reported to EH&S at x3194 within 8 hours.

Non-Life Threatening Emergency – Go to Student Health, building 588, x5361, Hours: M, T, R, F 8 a.m. to 4:30 p.m, W 9 a.m. to 4:30 pm and R 5 to 7 p.m. by appointment. At all other times report to Goleta Valley Cottage Hospital (emergency room) at 351 South Patterson Avenue, 805-967-3411. Contact PI and EH&S as soon as practical.Note: All serious injuries must be reported to EH&S at x3194 within 8 hours.

Decontamination/Waste Disposal Procedure

Wearing proper PPE, please decontaminate equipment and bench tops using water.  Please dispose of the used material and disposables contaminated with this material as hazardous waste.

General hazardous waste disposal guidelines:

Label Waste

- Affix a hazardous waste tag on all waste containers as soon as the first drop of waste is added to the container

Store Waste

- Store hazardous waste in closed containers, in secondary containment and in a designated location

- Double-bag dry waste using transparent bags

- Waste must be under the control of the person generating & disposing of it

Dispose of Waste

- Dispose of regularly generated chemical waste within 90 days

- Call EH&S for questions and for pick-up

- Empty Containers: Dispose as hazardous waste irrespective of the container size

Prepare for transport to pick-up location

- Check waste tag

- Use secondary containment

Safety Data Sheet (SDS) Location

Online SDS can be accessed at http://www.chem.ucsb.edu/about/safety .

Protocol/Procedure

Electrophoresis involves application of voltage across an aqueous gel. The levels we use are typically in the ‘hazardous’ range. Always use the provided plexiglass covers for the gel so that nothing can fall into the apparatus or cause a short circuit. Turn off power before manipulating any part of the apparatus. Check for leaks – if buffer drains out of one compartment, sparks can fly across the voltage difference. Most power supplies have the ability to cut off power if the current or voltage go above a threshold value – use this safety feature. If a wire breaks, replace the wire rather than attempt to repair it.

General electrical work:

- Make sure equipment is both off and unplugged.

- Make connections properly, that are both physically strong and well connected electrically.  Screwing down a wire to a metal connection, or soldering are good ways to do this.

- Do not leave sloppy strands of wire sticking off of any connection

- Cover any exposed wire with heat-shrink tubing, or allow the wire insulation to end exactly at the connection.  If the connection is outside of the instrument casing (ex: connecting 2 wires that stretch across a desk), cover the connection with heat shrink tubing, or another insulating shield.

- Check that no shorts were created by using a multimeter to measure the resistance from ground to line (>1MΩ), before turning equipment back on.

Soldering

- Thoroughly clean all surfaces to be soldered, and roughen them using a heavy grit sand paper.

- Make the connection first (before soldering) so that it is physically stable.  This usually involves clamping or wrapping an auxiliary wire around the 2 spliced wires.  NASA splicing guidelines are good guides for how to do this for different connections.  Set up heat shrink tubing so that it can be slid on after soldering.

- Use rosin-core soldering.  It makes everything easier.  Do not use extra flux.

- Under an elephant trunk hood, heat up the soldering iron, and dampen a piece of sponge or paper towel.  When the gun is hot, use solder to clean the iron, wiping off dirty solder with the damp utensil.  The paper towel or sponge should NOT burn or melt if it is damp.  The idea is that the water is absorbing the heat.

- When the iron is clean, put a drop of solder on the tip, and hold the tip to the joint to be soldered.  Hold the solder on the other side.  The drop on the tip is for heat transfer.  When the joint gets hot, the solder on the other side will melt and flow over the joint.  This makes a better connection than if solder is simply dripped or wiped onto a joint that is not being wet well by the solder.

- If it does not work, roughen the surface well and try again.  The iron may not be hot enough.

- The hardened flux will chip off easily.  For stubborn flux, a razor blade works well.

- Cover exposed metal with heat-shrink tubing, etc.

Documentation of Training (signature of all users is required)

I have read and understand the content of this SOP:

Name                                             Signature                                      Date